THERMATEST Model TCFCM Comparative Instrument

Thermal conductivity-the rate of heat flow through a substance-has become a parameter of critical importance in many fields and applications. In materials development for electronics  packaging  applications, for example, the ability of a material to conduct heat is the governing factor in determining its performance.

The Model TCFCM comparative instrument is a convenient, flexible system which mesures the thermal conductivity of a variety of solid materials over a broad range of  temperatures and environmental conditions. Materials   wich   can be tested in the   model TCFCM-N20, include ceramics, composites, metals, metals alloys, filled plastics and epoxies, geological materials, and carbon products.

This instruments an excellent research tool for evaluating materials in the moderate-to high conductivity range of 0.2 to 100   W/m K. Standard  sample sizes are 25 or 50mm in diameter and 12 to 38mm thick. Test can be performed over a temperatures range of -150ºC to +1000ºC, either in air or inert gas environments. Liquid nitrogen  must  be used to allow operation at cryogenic temperatures.

The comparative technique utilizes  known   propieties of standard reference materials to measure heat flow. The test specimen is sandwiched between two identical  reference samples. This stack is placed   beetween two heating elements, controlled at different temperatures. The resultant flow of heat  through  the composite stack produces  a temperature distribution wich is measured  with termocouples. At thermal equilibrium, thermal conductivity of the test specimen can then be computed.

This instrument is designed to archieve test accuracies between +-5% to+-10% over the entire thermal conductivity range. An important factor in estimating the overall test accuracy  is the extent  to wich  the thermal conductance of the test sample matches that of the reference material . Several types of reference materials  are offered  wich cover different segments of thermal conductivity range, including  pyrex 7740, pyroceram 9606, inconel 718, and electrolytic iron.

A cylindrical linear gradient guard heater is placed around the samples to ensure that the heat  flux through the stack is constant .The top and bottom temperatures of the  guard heater are controlled  at the same  set points as the corresponding heaters  in the stack, thus producing an axial gradient in the guard  heater to that in the test stack. Temperatures along the length of the guard heater can be measured to verify proper control.

### Specifications-  Model  TCFCM

| Test Method | Comparative Thermal Conductivity |
| --- | --- |
| Standard Sample Sizes | 50 mm diameter 25mm diameter 12 a 38 mm Thick |
| Temperature Range | -150ºC a +1000ºC |
| Sample Environment | Air, inert gas |
| Accuracy | (5% a (10% |
| Reproducibility | (3% |
| Temperature Sensors | Chromel/Alumel Thermocouples |
| Heat Sink Coolant | T ( 100ºC: city water T(100ºC : Opcional circulator or LN2 |

Services:

To mesure   thermal conductivity of ceramics, composites, metals, metals alloys, filled plastics and epoxies, geological materials, and carbon products.

Contact :

M.Núria Vives

Tel.+ 34 93 401 59 00

Fax+34 93 401 73 89

e-mail: nvives@mmt.upc.edu

Adress:

Unidat  de Termotècnia

Escola Técnica Superior d’Ingeniería Industrial de  Barcelona

(Universidat  Politécnica de Catalunya)

Avinguda diagonal nº 647

08028 Barcelona

Spain